## Slide 1
PRINCIPLES OF FIELD SPECTROSCOPY • GEO-MSc-704
MODULE 3
Sensor Calibration, White Reference (Spectralon) & Radiometric Validation
4 Teaching Hours • 1.5 ECTS Credits • 30 Lecture Slides
Module 3 details laboratory metrological calibration and in-field radiometric validation procedures for field spectroradiometers. Students examine NIST-traceable standards, atomic emission line lamps (Hg-Ar) and polymer absorption standards (Mylar) for sub-nanometer wavelength calibration, and large integrating spheres for absolute spectral radiance responsivity mapping. The module investigates Spectralon sintered PTFE white reference physics, non-Lambertian departures at high solar zenith angles, degradation and cleaning protocols, cosine receptor characterization, and the construction of comprehensive measurement uncertainty budgets conforming to the ISO/BIPM Guide to the Expression of Uncertainty in Measurement (GUM).

## Slide 2
MOD 3 • [QUALITYASSURANCE] • SLIDE 1/30
Principles of Metrological Traceability in Spectroscopy
CORE TAKEAWAY: Traceability establishes an unbroken chain of calibrated comparisons connecting field measurements back to international SI standards (NIST/NPL).
Metrological traceability: Property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty.
Primary standards: National Metrology Institutes (e.g. NIST in USA, NPL in UK, PTB in Germany).
Secondary/Transfer standards: Laboratory integrating spheres, calibrated FEL tungsten lamps, and certified Spectralon panels.
FIELD PROTOCOL NOTE
Without NIST/NPL traceability, ground validation spectra cannot be accepted for satellite Cal/Val programs.
References: JCGM (2008). International Vocabulary of Metrology (VIM). Joint Committee for Guides in Metrology.
Notes: Welcome to Module 3. We focus on sensor calibration, white reference integrity, and rigorous uncertainty propagation.

## Slide 3
MOD 3 • [INSTRUMENTATION] • SLIDE 2/30
Wavelength Calibration: Physics of Gas Emission Lines
CORE TAKEAWAY: Wavelength calibration maps detector channel index numbers to absolute physical wavelengths using sharp atomic emission lines.
Gas discharge pencil lamps: Low-pressure Mercury-Argon (Hg-Ar), Neon, Krypton, or Xenon.
Atoms excited by electric discharge emit photons at invariant, quantum-mechanically defined transition wavelengths.
Mercury (Hg) primary lines: 404.66 nm, 435.84 nm, 546.07 nm, 576.96 nm, 579.07 nm.
Argon (Ar) lines: Dense emission line clusters in the 690–900 nm region.
MATHEMATICAL FORMULATION
\lambda_c = a_0 + a_1 \cdot x + a_2 \cdot x^2 + a_3 \cdot x^3
Polynomial dispersion function relating pixel position x to calibrated wavelength λ_c.
References: Reader, J. et al. (1980). Wavelengths and Transition Probabilities for Atoms and Atomic Ions. NIST.
Notes: Show the emission spectrum of a Hg-Ar lamp. Emphasize that these physical atomic transitions never drift over time.

## Slide 4
MOD 3 • [INSTRUMENTATION] • SLIDE 3/30
Wavelength Calibration in the SWIR: Mylar & Rare Earth Filters
CORE TAKEAWAY: Solid rare-earth oxide and polymer standards provide narrow absorption lines in the SWIR (1000–2500 nm) where gas lamps have sparse lines.
Holmium Oxide (Ho₂O₃) and Dysprosium Oxide: Certified absorption bands in the NIR and SWIR.
Mylar film (polyester): Sharp C-H and ester vibrational absorption spikes at 1660 nm, 2135 nm, 2260 nm, and 2470 nm.
Band-centroid fitting: Fitting a parabolic or Gaussian peak over the observed absorption minimum yields wavelength accuracy < 0.2 nm.
FIELD PROTOCOL NOTE
Check wavelength stability annually or following any physical drop or high-vibration vehicle transport.
References: MacArthur, A. et al. (2012). Calibration and characterization of field spectrometers. NERC FSF.
Notes: Describe the practical procedure for placing a Mylar film over the entrance foreoptic under a stable halogen lamp.

## Slide 5
MOD 3 • [INSTRUMENTATION] • SLIDE 4/30
Absolute Radiometric Calibration: Integrating Spheres
CORE TAKEAWAY: Radiometric calibration determines absolute spectral responsivity, converting raw digital counts into physical radiance [W m⁻² sr⁻¹ nm⁻¹].
Calibration source: Large internally illuminated integrating sphere (e.g. 50 cm or 100 cm diameter).
Illumination: Multiple regulated quartz-tungsten-halogen (QTH) lamps operating with power supplies stable to ±0.01% current.
Exit port: Uniform diffuse radiance field across the entire exit aperture (spatial uniformity > 99%).
MATHEMATICAL FORMULATION
R_{\text{cal}}(\lambda) = \frac{\text{DN}_{\text{sphere}}(\lambda) - \text{DN}_{\text{dark}}(\lambda)}{t_{\text{int}} \cdot L_{\text{sphere}}(\lambda)}
Instrument spectral responsivity calibration factor R_cal [counts / (W m⁻² sr⁻¹ nm⁻¹ · ms)].
References: Walker, J. H. et al. (1987). Spectral Radiance Calibrations. NBS Special Publication 250-1.
Notes: Explain why integrating spheres must be calibrated against NIST gold-standard blackbody cavity furnaces.

## Slide 6
MOD 3 • [INSTRUMENTATION] • SLIDE 5/30
Irradiance Calibration: Calibrated Standard FEL Lamps
CORE TAKEAWAY: Standard 1000-Watt FEL tungsten-halogen lamps provide calibrated spectral irradiance at an exact distance of 50.0 cm.
FEL lamps operate at ~3100 K color temperature inside a dark baffle tunnel to eliminate stray ambient wall reflections.
Inverse-square law: Spectral irradiance scales inversely with the square of the distance from the lamp filament plane.
Used to calibrate cosine receptor diffusers for downwelling solar irradiance measurement.
MATHEMATICAL FORMULATION
E(\lambda, r_2) = E(\lambda, r_1) \cdot \left(\frac{r_1}{r_2}\right)^2
Inverse-square distance relationship for point-source irradiance calibration.
References: Yoon, H. W. et al. (2003). Realization of the NIST detector-based spectral irradiance scale. Metrologia.
Notes: Point out that a 1 mm positioning error in lamp distance at 50 cm produces a 0.4% calibration error.

## Slide 7
MOD 3 • [QUALITYASSURANCE] • SLIDE 6/30
Dark Current Compensation: Methods and Drift
CORE TAKEAWAY: Dark current must be recorded with an optical shutter closed and subtracted from raw digital counts prior to any mathematical operations.
Electronic baseline: Even with zero incident photons, thermal energy excites electrons into detector wells.
Integration time scaling: Dark current counts accumulate linearly with exposure time in silicon arrays.
Dynamic drift: Internal instrument heating causes dark current to drift continuously during outdoor field campaigns.
FIELD PROTOCOL NOTE
Always re-acquire dark current every 10–15 minutes, or whenever ambient temperature shifts noticeably.
References: MacArthur, A. (2007). A field procedure to compensate for temperature drift. NERC FSF.
Notes: Emphasize that dark current subtraction is non-negotiable; forgetting it renders all reflectance math invalid.

## Slide 8
MOD 3 • [INSTRUMENTATION] • SLIDE 7/30
Non-Linearity Correction of Detector Electronics
CORE TAKEAWAY: Real photodiode arrays exhibit non-linear count responses as charge approaches the full well capacity of the junction.
Ideal linear detector: Double the photon flux produces exactly double the digital numbers (DN).
At high charge levels (> 80% full well), recombination rates increase, causing sub-linear response.
Characterized using beam-addition techniques or calibrated flux attenuator wheels.
Non-linearity correction polynomials linearize digital counts to < 0.1% departure across the dynamic range.
MATHEMATICAL FORMULATION
\text{DN}_{\text{linear}} = \sum_{k=0}^p c_k \cdot (\text{DN}_{\text{raw}} - \text{DN}_{\text{dark}})^k
Polynomial linearization function mapping raw counts to idealized linear response.
References: Saunders, R. D., & Shumaker, J. B. (1984). Automated radiometric linearity tester. Applied Optics.
Notes: Show the departure from linearity at counts > 55,000 on a 16-bit ADC, explaining why 45,000 counts is the target optimization limit.

## Slide 9
MOD 3 • [QUALITYASSURANCE] • SLIDE 8/30
Spectralon® White Reference Reflectance Calibration Certificate
CORE TAKEAWAY: No white reference panel is 100% reflective; NIST calibration certificates provide exact spectral correction factors (98–99%).
A panel assumed to be 100% introduces a systematic 1% to 2% underestimation of true target reflectance if uncorrected.
Calibration certificates provide 8°/hemispherical reflectance factors at 50 nm intervals from 250 to 2500 nm.
Interpolation: Cubic spline interpolation maps certificate values onto the 1 nm instrument channels.
MATHEMATICAL FORMULATION
\rho_{\text{target}}(\lambda) = \frac{\text{DN}_{\text{target}}(\lambda) - \text{DN}_{\text{dark}}(\lambda)}{\text{DN}_{\text{ref}}(\lambda) - \text{DN}_{\text{dark}}(\lambda)} \cdot R_{\text{certificate}}(\lambda)
Proper application of the NIST-traceable Spectralon calibration vector R_certificate.
FIELD PROTOCOL NOTE
Store the active calibration certificate file in the field laptop and verify matching panel serial numbers.
References: Labsphere Inc. (2018). Spectralon Reflectance Standards Technical Guide.
Notes: Walk students through reading a genuine Labsphere calibration certificate and extracting the tabular data.

## Slide 10
MOD 3 • [QUALITYASSURANCE] • SLIDE 9/30
Spectralon Degradation: UV Solarization & Soil Dust Contamination
CORE TAKEAWAY: Field use inevitably degrades Spectralon panels; annual recalibration is essential for data credibility.
UV Solarization: Extended exposure to intense outdoor solar UV radiation breaks polymer bonds, causing yellowing and UV absorption.
Dust and Organic Particulates: Windblown fine dust lodges permanently into the microscopic porous PTFE matrix.
Hydrocarbon Stains: Touching the panel with oily fingers introduces C-H absorption troughs in the SWIR (1200, 1720, 2300 nm).
FIELD PROTOCOL NOTE
Never blow dust off a panel with your mouth (saliva droplets ruin the surface); use clean dry nitrogen or a photographic air bulb.
References: MacArthur, A. et al. (2012). NERC FSF Technical Report.
Notes: Show before-and-after spectral curves of an abused field panel versus a pristine reference.

## Slide 11
MOD 3 • [QUALITYASSURANCE] • SLIDE 10/30
Spectralon Resurfacing & Cleaning Protocol
CORE TAKEAWAY: Spectralon panels can be restored to near-factory reflectance by gentle wet-sanding under deionized water.
Equipment: Waterproof 220–400 grit silicon carbide emery sandpaper, glass plate, deionized water.
Procedure: Place sandpaper on flat glass, wet thoroughly, sand panel in a figure-eight pattern under gentle pressure.
Bakeout: Dry in a clean vacuum oven at 80°C for 8–12 hours to drive off residual water and volatile organics.
Post-sanding: Must be recalibrated or cross-calibrated against a secondary master laboratory panel.
FIELD PROTOCOL NOTE
Perform cleaning only in a clean laboratory environment, never in the field hotel or tent.
References: Labsphere Inc. (2015). Care and Handling of Spectralon Diffuse Reflectance Targets.
Notes: Describe the Labsphere recommended sanding procedure step-by-step.

## Slide 12
MOD 3 • [QUALITYASSURANCE] • SLIDE 11/30
Secondary Field Panels vs. Master Laboratory Standards
CORE TAKEAWAY: Protect master calibration standards by using dedicated secondary panels in the field and reserving pristine masters for lab checks.
Golden Master Panel: Kept permanently in a dark, desiccated laboratory cabinet, never taken into the field.
Field Working Panels: Deployed in outdoor campaigns and subjected to dust, sun, and transport.
Cross-calibration: Measure field working panels against the golden master in the lab before and after every field campaign.
Tracking drift: Generates an empirical degradation vector that can retroactively correct field data.
References: Milton, E. J. et al. (2009). Progress in field spectroscopy. Remote Sensing of Environment.
Notes: Instill this institutional discipline: field working standards will get dirty; master standards must remain untouched.

## Slide 13
MOD 3 • [PHYSICS] • SLIDE 12/30
White Reference Non-Lambertian Behavior at High Zenith Angles
CORE TAKEAWAY: Spectralon is not perfectly Lambertian; its reflectance factor increases significantly when solar zenith angles exceed 50°–60°.
At normal incidence (θ_i < 30°), Spectralon is isotropic within ±0.5%.
At grazing angles (θ_i > 60°), forward specular scattering emerges, and measured apparent reflectance rises by 2% to 6%.
If an operator uses a Spectralon panel at 16:00 in winter (θ_s = 65°), target reflectance is artificially suppressed by up to 5%.
MATHEMATICAL FORMULATION
R(\theta_i, 0; \text{nadir}) = 1.0 + c_1 \cdot \theta_i^2 + c_2 \cdot \theta_i^4
Empirical model for non-Lambertian Spectralon departure as a function of solar incidence angle.
References: Sandmeier, S. et al. (1998). Physical mechanisms in hyperspectral BRDF data of grass and watercress. Remote Sensing of Environment.
Notes: Show the BRDF scattering polar plot of Spectralon measured by sandmeier et al. using a laboratory goniometer.

## Slide 14
MOD 3 • [INSTRUMENTATION] • SLIDE 13/30
Cosine Receptor Characterization: Directional Response Error
CORE TAKEAWAY: Cosine diffusers measure irradiance by capturing photons according to Lambert's cosine law; departures cause cosine error.
Ideal cosine receptor: Responsivity scales perfectly with cos(θ_i).
Real PTFE diffusers: Under-respond at high zenith angles (θ_i > 60°) due to surface reflection and rim shading.
Cosine error: f₂(θ) = (Signal(θ) / [Signal(0) · cos θ]) - 1.0.
High-grade cosine receptors maintain total cosine error f₂ < 3% up to 70° zenith angle.
MATHEMATICAL FORMULATION
f_2 = \int_0^{\pi/2} \left| \frac{S(\theta)}{S(0)\cos\theta} - 1 \right| \sin(2\theta) \, d\theta
CIE directional response index quantifying total cosine error.
References: CIE (1987). Methods of Characterizing the Performance of Radiometers and Photometers. Publication CIE 69.
Notes: Show the polar response curve of a cosine diffuser against the theoretical cosine circle.

## Slide 15
MOD 3 • [QUALITYASSURANCE] • SLIDE 14/30
Single-Beam Time-Drift Uncertainty
CORE TAKEAWAY: In single-beam field radiometry, the time elapsed between white reference and target scans is a primary uncertainty driver.
Atmospheric optical depth and solar position change continuously.
At solar noon in summer, solar zenith angle changes slowly (< 0.1° per minute).
In mid-morning or late afternoon, solar zenith angle changes by > 0.5° per minute, altering irradiance by ~1% every 2 minutes.
Rule of thumb: Complete target measurements within 3 to 5 minutes of white reference acquisition under clear skies.
FIELD PROTOCOL NOTE
Set an automatic timer on the field computer to alert the team when 5 minutes have elapsed since the last WR scan.
References: Milton, E. J. (1987). Principles of field spectroscopy. International Journal of Remote Sensing.
Notes: Demonstrate how an uncorrected 5-minute delay creates a false spectral slope across the entire measurement.

## Slide 16
MOD 3 • [INSTRUMENTATION] • SLIDE 15/30
Dual-Beam Inter-Calibration Protocols
CORE TAKEAWAY: Dual-beam systems require frequent cross-calibration of the reference and target sensors to track inter-channel drift.
Dual-beam eliminates solar temporal drift by recording WR and target simultaneously.
However, if the two spectrometers drift differentially in temperature or responsivity, a systematic cross-channel error occurs.
Protocol: Every 30–60 minutes, place both optical heads over the same white reference panel to calculate the cross-calibration vector G(λ).
MATHEMATICAL FORMULATION
G(\lambda) = \frac{L_1^{\text{WR}}(\lambda)}{L_2^{\text{WR}}(\lambda)}, \quad \rho_{\text{target}}(\lambda) = \frac{L_1^{\text{target}}(\lambda)}{L_2^{\text{WR}}(\lambda) \cdot G(\lambda)} \cdot \rho_{\text{ref}}(\lambda)
Cross-calibration gain vector G(λ) continuously normalizes inter-channel differences.
References: MacArthur, A. et al. (2014). Remote Sensing of Environment.
Notes: Emphasize that dual-beam is not magic: it trades temporal solar uncertainty for inter-channel calibration uncertainty.

## Slide 17
MOD 3 • [QUALITYASSURANCE] • SLIDE 16/30
GUM Framework: Guide to the Expression of Uncertainty in Measurement
CORE TAKEAWAY: Scientific measurements must be accompanied by an uncertainty statement compliant with international metrological standards.
JCGM 100:2008 (GUM): Defines standardized methodologies for evaluating and expressing measurement uncertainty.
Type A Evaluation: Evaluated by statistical analysis of series of observations (e.g. sample standard deviation of repeated scans).
Type B Evaluation: Evaluated by scientific judgment, manufacturer certificates, calibration standards, or published literature.
MATHEMATICAL FORMULATION
u_c^2(y) = \sum_{i=1}^N \left(\frac{\partial f}{\partial x_i}\right)^2 u^2(x_i) + 2\sum_{i=1}^{N-1}\sum_{j=i+1}^N \frac{\partial f}{\partial x_i}\frac{\partial f}{\partial x_j} u(x_i, x_j)
Law of propagation of uncertainty for combined standard uncertainty u_c(y).
References: JCGM (2008). Evaluation of measurement data — Guide to the expression of uncertainty in measurement. JCGM 100:2008.
Notes: Introduce the GUM methodology. Remote sensing papers that state "reflectance = 0.42" without ± uncertainty are methodologically deficient.

## Slide 18
MOD 3 • [QUALITYASSURANCE] • SLIDE 17/30
Building a Field Spectroscopy Uncertainty Budget
CORE TAKEAWAY: A complete uncertainty budget aggregates all independent noise and systematic error components into a combined uncertainty curve.
Component 1: Spectralon certificate uncertainty (u_ref ≈ 0.5%).
Component 2: White reference aging/contamination (u_panel ≈ 1.0%).
Component 3: Instrument noise / SNR (u_noise ≈ 0.2% in VNIR, 1.5% in SWIR).
Component 4: Solar irradiance temporal drift (u_sun ≈ 0.5–2.0%).
Component 5: Zenith/nadir pointing tilt error (u_tilt ≈ 0.8%).
FIELD PROTOCOL NOTE
Total combined relative reflectance uncertainty typically ranges from 1.5% to 3.0% under optimal clear-sky conditions.
References: Fox, N. et al. (2003). Traceability in remote sensing: The QE6 project. SPIE Proceedings.
Notes: Walk students through a spreadsheet table totaling the individual variance terms in quadrature.

## Slide 19
MOD 3 • [MATHEMATICS] • SLIDE 18/30
Wavelength Uncertainty vs. Band-Depth Retrieval Error
CORE TAKEAWAY: A tiny 1 nm wavelength shift creates massive errors when calculating narrow absorption band depths in minerals and pigments.
Kaolinite Al-OH doublet at 2206 nm has a half-width of only ~12 nm.
A 1.0 nm wavelength calibration error produces an apparent 5% to 10% error in calculated band absorption depth.
In vegetation, red edge inflection point (REIP) shifts of 2–3 nm indicate heavy metal stress or nitrogen deficiency; wavelength stability < 0.5 nm is mandatory.
References: Van der Meer, F. (2004). Analysis of spectral absorption features in hyperspectral imagery. International Journal of Applied Earth Observation.
Notes: Show the derivative of a sharp absorption peak: slope is steepest on the wings, where wavelength shifts generate maximum error.

## Slide 20
MOD 3 • [INSTRUMENTATION] • SLIDE 19/30
Radiometric Calibration Transfer Standards: Portable Field Lamps
CORE TAKEAWAY: Portable field calibration lamps allow on-site verification of instrument responsivity during remote field expeditions.
Compact stabilized QTH light sources with precision DC constant-current power regulation.
Mounted directly to the foreoptic barrel to provide a repeatable radiant flux baseline.
Enables detection of sudden optical component failure (e.g. fractured fiber strand or misaligned order-sorting filter) while still in the field.
FIELD PROTOCOL NOTE
Run a portable lamp check at the field base camp every morning before setting off to remote field sites.
References: NERC FSF (2020). Calibration Equipment and Field Validation Manual.
Notes: Describe the NERC FSF field validation lamp kits.

## Slide 21
MOD 3 • [QUALITYASSURANCE] • SLIDE 20/30
Inter-Comparison Campaigns & Round-Robin Experiments
CORE TAKEAWAY: Round-robin experiments comparing multiple spectrometers over identical targets reveal inter-instrument biases of 2% to 8%.
ESA and CEOS regularly conduct field spectroradiometer inter-comparison campaigns (e.g. RadCalNet, FICE).
Different instruments (ASD vs SVC vs PSR) measuring the same Spectralon panel simultaneously frequently differ by 2–4% in the SWIR.
Major causes: Differing slit functions, stray light rejection, fiber cable geometries, and manufacturer calibration algorithms.
References: Anderson, K. et al. (2011). Field spectroradiometer inter-calibration: Methods and challenges. Remote Sensing of Environment.
Notes: Discuss results from the CEOS WGCV inter-comparison campaigns over dry lake salt flats.

## Slide 22
MOD 3 • [QUALITYASSURANCE] • SLIDE 21/30
CEOS WGCV Field Spectrometry Protocol
CORE TAKEAWAY: The Committee on Earth Observation Satellites defines the global standard for ground validation of spaceborne missions.
Protocol requirements: Traceable calibration certificates < 1 year old; clear sky criteria (AOD < 0.2); SZA < 50°.
Sampling footprint documentation: Drone imagery of target site, high-precision RTK GNSS coordinates (±2 cm).
Uncertainty statement: Every published spectral curve must report a combined standard uncertainty vector u(λ).
FIELD PROTOCOL NOTE
Review the CEOS WGCV checklist prior to any satellite under-flight field validation campaign.
References: CEOS WGCV (2014). Quality Assurance Framework for Earth Observation: Field Spectrometry Protocol.
Notes: Show the official CEOS document and review its mandatory field log parameters.

## Slide 23
MOD 3 • [CASESTUDY] • SLIDE 22/30
RadCalNet: Automated Ground Calibration Networks
CORE TAKEAWAY: RadCalNet provides automated, continuous in situ surface reflectance measurements for autonomous satellite vicarious calibration.
Global network of automated instrumented sites: Railroad Valley (USA), La Crau (France), Baotou (China), Gobabeb (Namibia).
Autonomous nadir-viewing multi-channel spectroradiometers taking measurements every 30 minutes.
Combined with co-located CIMEL sun photometers (AERONET) measuring atmospheric optical depth and water vapor column.
Provides top-of-atmosphere (TOA) and surface reflectance with documented uncertainty < 3% for Sentinel-2, Landsat-8/9, PlanetScope.
References: Bouvet, M. et al. (2019). RadCalNet: A radiometric calibration network for Earth observing imagers. Remote Sensing, 11(20), 2401.
Notes: Display the live RadCalNet web portal and inspect a real-time calibrated reflectance curve from Railroad Valley.

## Slide 24
MOD 3 • [CASESTUDY] • SLIDE 23/30
Vicarious Calibration of Earth Observation Satellites
CORE TAKEAWAY: Vicarious calibration uses ground field spectroscopy over bright, homogeneous targets to validate satellite onboard sensors.
Satellite sensors degrade in orbit due to cosmic radiation, solar UV darkening of lenses, and outgassing contamination.
Reflectance-based method: Measure ground surface reflectance with field spectroradiometers during satellite overpass.
Atmospheric radiative transfer: Input ground spectra and AERONET aerosol/gas data into MODTRAN to predict Top-of-Atmosphere (TOA) radiance.
Compare predicted TOA radiance against satellite digital counts to update sensor calibration gain coefficients.
MATHEMATICAL FORMULATION
L_{\text{TOA}}(\lambda) = L_{\text{path}}(\lambda) + \frac{\rho_{\text{ground}}(\lambda) \cdot E_{\text{solar}}(\lambda) \cdot T_{\downarrow}(\lambda) \cdot T_{\uparrow}(\lambda)}{\pi \cdot [1 - s(\lambda) \cdot \rho_{\text{ground}}(\lambda)]}
Simplified radiative transfer equation linking surface reflectance to satellite TOA radiance.
References: Thome, K. et al. (2004). Vicarious calibration of SPOT-4 HRVIR, Landsat-7 ETM+, and TERRA ASTER. IEEE TGRS.
Notes: Explain how field spectroscopy was critical for re-calibrating Landsat-7 ETM+ and Landsat-8 OLI after launch.

## Slide 25
MOD 3 • [QUALITYASSURANCE] • SLIDE 24/30
Spectrometer Field Validation Checklist
CORE TAKEAWAY: Rigorous pre-flight and pre-field checklist execution prevents costly data invalidation in retrospective analysis.
Phase 1 (Laboratory): Inspect fiber bundle, confirm wavelength alignment with Hg-Ar or Mylar, verify Peltier cooling lock.
Phase 2 (Departure): Clean Spectralon panels, charge all lithium batteries, verify calibration files loaded in software.
Phase 3 (On-Site): Level tripod, orient station to solar azimuth, run auto-integration, record 100% white reference check.
Phase 4 (Post-Campaign): Inspect raw data for detector splice jumps, saturation flags, and SNR degradation.
FIELD PROTOCOL NOTE
Laminate this 4-phase checklist and keep it permanently inside the spectrometer rugged transport case.
References: NERC FSF (2020). Operational Field Guide.
Notes: Distribute the laminated checklist template to students.

## Slide 26
MOD 3 • [PROCESSING] • SLIDE 25/30
Signal Clipping & Saturated Band Masking
CORE TAKEAWAY: Saturated channels distort spectral shape and create spurious absorption artifacts in downstream derivative analysis.
Saturation occurs when ADC counts hit maximum (e.g. 65,535 counts in 16-bit).
Partial saturation: Even at 63,000 counts, photodiode amplifiers may leave the linear regime.
Automated masking: Flag and set all spectral channels with DN > 0.90 · DN_max to NaN.
Never interpolate through saturated bands; flag them explicitly in the metadata record.
FIELD PROTOCOL NOTE
If bright cloud borders momentarily illuminate the panel during WR acquisition, the entire subsequent dataset is ruined.
References: Analytical Spectral Devices (2012). RS3 User Guide.
Notes: Show a spectrum where one channel was saturated and how it created a massive spike when computing NDVI.

## Slide 27
MOD 3 • [MATHEMATICS] • SLIDE 26/30
Spectral Interpolation: Splines vs. Linear vs. PCHIP
CORE TAKEAWAY: Interpolating irregular or missing spectral channels requires shape-preserving algorithms to prevent artificial oscillations.
Linear interpolation: Fast and robust, but introduces sharp non-physical slope breaks at tie points.
Cubic spline: Smooth curves, but susceptible to Runge's phenomenon (overshooting and spurious wiggles) near steep absorption edges.
PCHIP (Piecewise Cubic Hermite Interpolating Polynomial): Preserves monotonicity and avoids artificial overshoot.
MATHEMATICAL FORMULATION
y(x) = h_{00}(t)y_k + h_{10}(t)d_k + h_{01}(t)y_{k+1} + h_{11}(t)d_{k+1}
Hermite basis formulation preserving monotonic derivative transitions.
References: Fritsch, F. N., & Carlson, R. E. (1980). Monotone piecewise cubic interpolation. SIAM Journal on Numerical Analysis.
Notes: Demonstrate how a cubic spline creates a false absorption feature near the steep red edge if tie points are sparse.

## Slide 28
MOD 3 • [PROCESSING] • SLIDE 27/30
Atmospheric Noise Truncation Protocols
CORE TAKEAWAY: Channels in deep atmospheric water vapor absorption bands (1350–1450 nm and 1800–1980 nm) must be masked out.
Zero incident solar irradiance means dividing target counts by zero or near-zero reference counts.
Produces extreme random division spikes with noise reaching ±1000% reflectance.
Academic standard: Truncate and mask out: 1350–1460 nm and 1790–1990 nm prior to publication or modeling.
Truncate noisy ends below 380 nm and above 2450 nm where detector responsivity is low.
FIELD PROTOCOL NOTE
Never present field reflectance curves showing noisy spiky grass in the 1400 nm and 1900 nm troughs.
References: Milton, E. J. et al. (2009). Remote Sensing of Environment.
Notes: Explain why keeping these water vapor bands in statistical models (PLSR/Random Forest) leads to catastrophic overfitting.

## Slide 29
MOD 3 • [CASESTUDY] • SLIDE 28/30
Laboratory Cross-Calibration Experiment Design
CORE TAKEAWAY: A structured laboratory experiment determines the relative spectral responsivity ratio between two field instruments.
Set up a stabilized 1000W QTH lamp illuminating a pristine 10" master Spectralon panel.
Mount both instruments on identical viewing jigs with optical axes aligned at 45° incidence, 0° nadir viewing.
Acquire 50 scans per instrument, swap physical positions, and repeat to eliminate spatial illumination gradients.
Calculates cross-instrument transfer matrix with relative standard uncertainty < 0.3%.
References: MacArthur, A. et al. (2012). NERC FSF Technical Report.
Notes: Present the experimental layout diagram and show student-collected cross-calibration curves.

## Slide 30
MOD 3 • [QUALITYASSURANCE] • SLIDE 29/30
Calibration Record Keeping and Audit Trails
CORE TAKEAWAY: Reproducible geospatial science requires permanent digital audit trails linking raw scans to specific calibration files.
Metadata requirements: Spectrometer serial number, fiber optic bundle ID, foreoptic serial number.
White reference panel serial number and calibration certificate issue date.
Internal instrument operating firmware version and dark current subtraction timestamp.
Without these records, historic field data cannot be re-calibrated when standards improve.
FIELD PROTOCOL NOTE
Store calibration files alongside raw spectral files in read-only data repositories.
References: Wilkinson, M. D. et al. (2016). The FAIR Guiding Principles for scientific data management. Scientific Data, 3, 160018.
Notes: Remind students of the FAIR data principles: Findable, Accessible, Interoperable, Reusable.

## Slide 31
MOD 3 • [QUALITYASSURANCE] • SLIDE 30/30
Module 3 Synthesis: Toward Rigorous Field Protocols
CORE TAKEAWAY: Calibration and uncertainty quantification transform raw digital numbers into defensible scientific measurements.
Module 3 synthesized: NIST traceability, wavelength validation (Hg-Ar, Mylar), radiometric calibration, Spectralon care, and GUM uncertainty budgets.
In Module 4, we step outdoors into dynamic environments: solar zenith windows, BRDF sampling geometry, footprint calculations, and field metadata.
Laboratory practical: Computing GUM uncertainty propagation curves and building automated Spectralon contamination detectors.
References: Milton, E. J. et al. (2009). Progress in field spectroscopy. Remote Sensing of Environment, 113, S92-S109.
Notes: Conclude calibration module, invite questions on the uncertainty budget, and preview Field Exercise 3.