## Slide 1
PRINCIPLES OF FIELD SPECTROSCOPY • GEO-MSc-704
MODULE 2
Fundamentals of Point Spectroradiometers & Hardware Architecture
4 Teaching Hours • 1.5 ECTS Credits • 30 Lecture Slides
Module 2 deconstructs the internal physics, mechanical design, and electronic subsystems of commercial field spectroradiometers (ASD FieldSpec, SVC HR-1024, Spectral Evolution PSR+). Students analyze entrance foreoptics, fiber optic bundle numerical apertures, holographic diffraction gratings, order-sorting filters, and detector technologies spanning silicon photodiode arrays (VNIR) and thermoelectrically cooled InGaAs detectors (SWIR-1 and SWIR-2). The module details integration time optimization, dark current compensation, Spectralon reference standards, detector junction splice steps, and signal-to-noise ratio (SNR) modeling.

## Slide 2
MOD 2 • [INSTRUMENTATION] • SLIDE 1/30
Fundamentals of Point Spectroradiometers & Architecture
CORE TAKEAWAY: Point spectroradiometers integrate entrance optics, fiber guides, dispersive elements, and multi-detector arrays to record contiguous spectral signatures.
Point spectrometer: Measures spectral radiance or irradiance integrated over a single spatially averaged field of view.
Contrasts with imaging spectrometers (hyperspectral cameras) which collect 2D spatial pixels simultaneously.
High signal-to-noise ratio (SNR > 1000:1) and rapid integration times (10 to 100 ms) make point sensors the benchmark for ground validation.
FIELD PROTOCOL NOTE
Point spectrometers are precision laboratory-grade optical instruments packaged for ruggedized field deployment.
References: Curtiss, B., & Goetz, A. F. (1994). Field spectrometry: Techniques and instrumentation.
Notes: In Module 2, we study the physical hardware subsystems that convert incident photons into calibrated digital numbers.

## Slide 3
MOD 2 • [INSTRUMENTATION] • SLIDE 2/30
Anatomy of Benchmark Field Spectroradiometers
CORE TAKEAWAY: The three leading commercial field instruments (ASD FieldSpec, SVC HR-1024, Spectral Evolution PSR+) share core tripartite architecture.
ASD FieldSpec 4 Hi-Res: Tri-detector architecture (VNIR Si-array + SWIR1 InGaAs + SWIR2 InGaAs), fiber optic cable with 57 quartz fibers.
SVC HR-1024i: Internal high-resolution diffraction gratings, 1024 channels, dual internal cameras, GPS, and Bluetooth.
Spectral Evolution PSR-3500 / PSR+: Direct-mount foreoptics without fiber cables, reducing mechanical fiber bend noise.
References: MacArthur, A. et al. (2014). Optimizing the use of field spectrometers. Remote Sensing of Environment.
Notes: Show physical instrument cutaway diagrams and compare the internal spectrometer optical benches.

## Slide 4
MOD 2 • [INSTRUMENTATION] • SLIDE 3/30
The Entrance Optics: Foreoptics & Solid Angles
CORE TAKEAWAY: Foreoptics define the instrument's angular field of view (FOV), determining the spatial footprint and optical throughput.
Bare Fiber: Typically ~25° full-angle cone (acceptance angle governed by numerical aperture of quartz fiber).
Restrictor Lenses: 1°, 4°, 8°, 10° barrel lenses reduce the FOV for small targets or tall canopy towers.
Cosine Receptor: Opal glass or PTFE diffuser with 180° (2π steradian) hemispherical collection for measuring downwelling irradiance.
MATHEMATICAL FORMULATION
NA = n \cdot \sin\theta_a = \sqrt{n_{\text{core}}^2 - n_{\text{cladding}}^2}
Numerical aperture (NA) of fiber optic cable dictates maximum acceptance half-angle θ_a.
References: NERC FSF (2020). Best Practice Manual for Field Spectroradiometry.
Notes: Warn students: Attaching a 1° or 8° foreoptic lens changes optical throughput and requires a new white reference calibration.

## Slide 5
MOD 2 • [INSTRUMENTATION] • SLIDE 4/30
Fiber Optic Bundles: Transmission & Care
CORE TAKEAWAY: Fiber optic cables are brittle silica glass bundles vulnerable to micro-fractures, bend-loss, and mode-scrambling defects.
Fiber construction: High-purity anhydrous fused silica (low-OH) core to minimize internal hydroxyl absorption at 1400 nm.
Randomized bundle: Transmits light from the circular entrance slit to three separate internal spectrometer entrance slits.
Mechanical bend attenuation: Bending the cable beyond minimum bend radius (> 15 cm) induces transient transmission loss and calibration drift.
FIELD PROTOCOL NOTE
Never step on, twist, or clamp a spectrometer fiber cable. Secure the cable along the pistol grip with hook-and-loop straps.
References: Milton, E. J. (1987). Principles of field spectroscopy. International Journal of Remote Sensing.
Notes: Demonstrate the fragile inner structure of a fiber cable. Broken strands permanently reduce throughput and degrade calibration.

## Slide 6
MOD 2 • [INSTRUMENTATION] • SLIDE 5/30
Contact Probes and Plant Probes with Internal Halogen Sources
CORE TAKEAWAY: Contact probes eliminate outdoor solar fluctuations by providing constant artificial illumination under a sealed optical head.
Illumination: Low-noise quartz-tungsten-halogen (QTH) bulb operating at stabilized voltage.
Leaf clip / Plant probe: Holds leaves gently against low-reflectance black (or white) background backing without solar influence.
Advantage: Enables reliable spectral measurements under cloudy skies, indoors, or at night.
Limitation: Measures contact leaf optical properties, not canopy structure or BRDF.
FIELD PROTOCOL NOTE
The heat from high-intensity halogen contact probes can cause rapid desiccation of delicate leaves within seconds.
References: Doughty, C. E. et al. (2011). Foliar spectral traits in tropical trees. Remote Sensing of Environment.
Notes: Always measure leaves with contact probe clamped for no more than 3–5 seconds to prevent thermal scorch.

## Slide 7
MOD 2 • [PHYSICS] • SLIDE 6/30
Dispersive Elements: Diffraction Gratings vs. Prisms
CORE TAKEAWAY: Reflective diffraction gratings disperse polychromatic light into angularly separated wavelengths via constructive interference.
Ruled vs. Holographic Gratings: Holographic gratings produce minimal stray light and eliminate ghost spectra.
Linear dispersion: Spatial separation of wavelengths along the detector plane: dx/dλ = f · m / (d · cos θ_m).
Diffraction orders: Order-sorting filters (long-pass or bandpass) are required to block overlapping second-order harmonics (e.g. 400 nm diffracted at 800 nm angle).
MATHEMATICAL FORMULATION
d \cdot (\sin\alpha + \sin\beta) = m \cdot \lambda
Grating equation where d is groove spacing, α is incident angle, β is diffracted angle, and m is spectral order.
References: Palmer, C. (2020). Diffraction Grating Handbook (8th ed.). Newport Corporation.
Notes: Explain how diffraction order overlap can cause false signals in the NIR if filter transitions degrade.

## Slide 8
MOD 2 • [INSTRUMENTATION] • SLIDE 7/30
Silicon Photodiode Arrays for VNIR Detection (350–1000 nm)
CORE TAKEAWAY: Silicon detectors operate on the internal photoelectric effect with high quantum efficiency in the visible and near-infrared.
512-element or 1024-element linear silicon NMOS/CMOS photodiode array.
Silicon bandgap energy Eg ≈ 1.12 eV at 300 K; limits long-wavelength detection cutoff to λ_cutoff = hc / Eg ≈ 1100 nm.
Fixed grating with no moving parts: All VNIR channels are exposed and read out simultaneously in milliseconds.
MATHEMATICAL FORMULATION
\lambda_{\text{cutoff}} = \frac{hc}{E_g} \approx \frac{1.24}{\text{Energy in eV}} \, [\mu\text{m}]
Bandgap cutoff limits the operational wavelength range of semiconductor detectors.
References: Sze, S. M. (2007). Physics of Semiconductor Devices. Wiley.
Notes: Show the quantum efficiency curve of silicon: peaks around 700–800 nm, drops sharply at 1000 nm.

## Slide 9
MOD 2 • [INSTRUMENTATION] • SLIDE 8/30
Indium Gallium Arsenide (InGaAs) Detectors for SWIR (1000–2500 nm)
CORE TAKEAWAY: InGaAs semiconductor alloys provide the sensitivity required to detect low-energy shortwave infrared photons.
Standard InGaAs (In₀.₅₃Ga₀.₄₇As lattice-matched to InP substrate): Covers SWIR-1 (1000–1700 nm).
Extended InGaAs: Graded ternary composition with lower bandgap (Eg ≈ 0.49 eV) covers SWIR-2 (1700–2500 nm).
Requires Thermoelectric (Peltier) cooling to -10°C to -25°C to suppress thermal dark current generation.
FIELD PROTOCOL NOTE
Allow the spectroradiometer 15–20 minutes warm-up time after powering on for the Peltier coolers to stabilize detector temperatures.
References: Rogalski, A. (2010). Infrared Detectors (2nd ed.). CRC Press.
Notes: Point out why older spectrometers used liquid nitrogen cooled detectors; modern thermoelectric coolers made portable field radiometry possible.

## Slide 10
MOD 2 • [INSTRUMENTATION] • SLIDE 9/30
Scanning Monochromator vs. Fixed Photodiode Array
CORE TAKEAWAY: Array spectrometers read all wavelengths concurrently; scanning grating instruments step sequentially with rotating galvanometers.
Array Spectrometer (e.g. ASD VNIR, SVC HR-1024): All spectral channels acquire simultaneously, freezing fast-moving cloud shadows.
Scanning Grating (e.g. ASD SWIR1 & SWIR2): A high-speed oscillating galvanometer scans the spectrum across a single thermoelectrically cooled InGaAs point detector.
Modern ASD FieldSpec 4 uses fast oscillating concave gratings: complete 350–2500 nm scan in 100 milliseconds.
References: Curtiss, B., & Goetz, A. F. (1994). Field spectrometry techniques.
Notes: Contrast the mechanical vulnerability of rotating scanning gratings with solid-state array designs.

## Slide 11
MOD 2 • [MATHEMATICS] • SLIDE 10/30
Spectral Resolution vs. Spectral Sampling Interval
CORE TAKEAWAY: Resolution (FWHM) measures optical band-separation power; sampling interval is the wavelength spacing between reported data points.
Full Width at Half Maximum (FWHM): Width of the instrument response function (bandpass) at 50% peak transmittance.
Sampling Interval: Wavelength step between adjacent digital output channels (e.g., interpolated to 1.0 nm).
Nyquist sampling theorem: To avoid undersampling, spectral sampling interval must be at least FWHM / 2.
MATHEMATICAL FORMULATION
\Delta \lambda_{\text{FWHM}} \ge 2 \cdot \delta\lambda_{\text{sampling}}
Criterion for resolving adjacent narrow spectral absorption features without aliasing.
References: CEOS WGCV (2014). Quality Assurance Framework for Earth Observation: Field Spectrometry Protocol.
Notes: Caution students: Just because a vendor exports data with 1 nm increments does not mean the optical resolution is 1 nm! ASD is 3 nm in VNIR, 8–10 nm in SWIR.

## Slide 12
MOD 2 • [PHYSICS] • SLIDE 11/30
Instrument Response Function (IRF) & Slit Function
CORE TAKEAWAY: The instrument slit function convolves monochromatic spectral lines into Gaussian-shaped instrument bandpasses.
Input monochromatic Dirac delta function δ(λ - λ₀) is mapped onto the detector as a Gaussian response function.
Determined experimentally using narrow gas emission lamps (Mercury-Argon, Neon, Krypton).
Essential for resynthesizing satellite spectral bands (e.g. Sentinel-2 MSI, Landsat-9 OLI) from continuous field spectra.
MATHEMATICAL FORMULATION
S(\lambda) = \int_{-\infty}^{+\infty} L_{\text{true}}(\lambda') \cdot f_{\text{slit}}(\lambda - \lambda') \, d\lambda'
Convolution integral describing the observed spectrum S(λ) through the spectrometer slit function.
References: Barsi, J. A. et al. (2014). The spectral response functions of Landsat-8 Operational Land Imager. Remote Sensing.
Notes: Walk students through the mathematical convolution of a hyperspectral curve with Sentinel-2 spectral response functions.

## Slide 13
MOD 2 • [INSTRUMENTATION] • SLIDE 12/30
Analog-to-Digital Conversion (ADC) & Dynamic Range
CORE TAKEAWAY: The ADC converts photo-generated analog voltage into discrete Digital Numbers (DN) with finite quantization resolution.
16-bit ADC provides 2¹⁶ = 65,536 discrete quantization levels (0 to 65535 counts).
Dynamic range must accommodate both bright 99% Spectralon panels in full sun and dark water bodies (< 2% reflectance).
Saturation: When incident photon flux generates voltage exceeding the maximum ADC reference voltage, counts peg at 65535.
MATHEMATICAL FORMULATION
\text{Dynamic Range [dB]} = 20 \cdot \log_{10}\left(\frac{\text{Full Well Capacity}}{\text{Readout Noise}}\right)
Defines the ratio between the maximum detectable signal and the noise floor.
FIELD PROTOCOL NOTE
Never record data with peak counts exceeding 85–90% of maximum ADC capacity to preserve headroom for solar bursts.
References: Holst, G. C. (2011). CCD Arrays, Cameras, and Displays. JCD Publishing.
Notes: Demonstrate what a clipped, flat-topped saturated spectrum looks like in software.

## Slide 14
MOD 2 • [PROTOCOL] • SLIDE 13/30
Integration Time Optimization
CORE TAKEAWAY: Integration time (exposure time) must be dynamically optimized on the white reference panel before scanning targets.
Integration time governs the duration charges accumulate on the photodiode capacitors before readout.
Too short: Low signal-to-noise ratio, dominated by quantization and readout noise.
Too long: Signal saturation (> 65535 counts), corrupting reflectance ratios.
Target optimum: Set integration time so peak white reference signal sits between 70% and 85% of saturation (~45,000 counts on 16-bit).
FIELD PROTOCOL NOTE
Run auto-integration optimization whenever solar zenith angle changes by > 5° or cloud conditions shift.
References: ASD Inc. (2010). FieldSpec 4 User Manual. Analytical Spectral Devices.
Notes: Show how integration time scales linearly with counts in VNIR array detectors.

## Slide 15
MOD 2 • [PROCESSING] • SLIDE 14/30
Detector Junction Steps / Jumps at 1000 nm & 1800 nm
CORE TAKEAWAY: Transitions between separate physical detectors create artificial discontinuities (splices) in raw spectral curves.
ASD FieldSpec employs three separate physical gratings/detectors: VNIR (350–1000 nm), SWIR-1 (1001–1800 nm), and SWIR-2 (1801–2500 nm).
Fiber strands feeding the three spectrometers have slight spatial misalignments, seeing slightly different target footprints.
Temperature shifts cause different dark current drifts between silicon and InGaAs arrays.
Produces characteristic "steps" or "jumps" at exactly 1000 nm and 1800 nm.
FIELD PROTOCOL NOTE
Splice correction algorithms mathematically adjust the SWIR1 and SWIR2 offsets to match the VNIR reference channel.
References: Dorigo, W. et al. (2006). A new sensor for vegetation reflectance measurements. IEEE TGRS.
Notes: Display an uncorrected vegetation spectrum showing a prominent 5% jump at 1000 nm. Explain splice correction.

## Slide 16
MOD 2 • [PHYSICS] • SLIDE 15/30
Signal-to-Noise Ratio (SNR) Physics
CORE TAKEAWAY: SNR quantifies spectral measurement precision; it is wavelength-dependent and peaks in the VNIR, dropping in atmospheric absorption bands.
SNR = Mean Signal (S) / Standard Deviation of Noise (σ_N).
Noise sources: Shot noise (Poisson statistics of photon arrival), Thermal dark noise, Readout noise, Digitization noise.
Shot noise scales with the square root of signal: SNR_shot = S / √S = √S.
MATHEMATICAL FORMULATION
\text{SNR}(\lambda) = \frac{\mu_S(\lambda)}{\sigma_N(\lambda)} \propto \sqrt{N_{\text{photons}}} \propto \sqrt{t_{\text{integration}} \cdot N_{\text{scans}}}
SNR improves with the square root of the number of averaged scans.
References: Bevington, P. R., & Robinson, D. K. (2003). Data Reduction and Error Analysis for the Physical Sciences. McGraw-Hill.
Notes: Explain why averaging 25 spectra improves SNR by a factor of 5 (√25 = 5), while averaging 100 scans improves it by 10.

## Slide 17
MOD 2 • [PROTOCOL] • SLIDE 16/30
Spectral Averaging in Field Environments
CORE TAKEAWAY: Averaging multiple consecutive scans suppresses random gaussian noise while preserving deterministic spectral signatures.
Standard operational protocol: Average 10 to 30 individual scans per stored sample spectrum.
Dark current averaging: 25 to 50 scans for a noise-free dark offset subtraction.
White reference averaging: 20 to 40 scans to ensure high-precision reference denominator.
Target averaging: 15 to 30 scans balanced against target movement and solar drift.
FIELD PROTOCOL NOTE
If wind causes canopy swaying, higher scan averaging blurs the signal across varying leaf angles.
References: NERC FSF (2020). Standard Operating Procedures.
Notes: Teach the rule of diminishing returns: moving from 25 to 100 scans quadruples measurement time but only doubles SNR.

## Slide 18
MOD 2 • [PHYSICS] • SLIDE 17/30
Dark Current: Physics and Generation Mechanisms
CORE TAKEAWAY: Dark current is thermal electron-hole pair generation in the semiconductor depletion region occurring in complete darkness.
Dark current is strongly temperature dependent, doubling approximately every 6°C to 8°C in silicon and InGaAs.
Dark current must be measured with an internal optical shutter closed and subtracted from every subsequent measurement.
Uncorrected dark current drift introduces a systematic positive offset that distorts low-reflectance targets (e.g. water, shadow).
MATHEMATICAL FORMULATION
I_{\text{dark}}(T) = A \cdot T^2 \cdot \exp\left(-\frac{E_g}{2 k_B T}\right)
Shockley-Read-Hall generation rate in semiconductor p-n junctions.
References: Sze, S. M. (2007). Physics of Semiconductor Devices.
Notes: Show the dark current count spectrum across the 2151 channels, highlighting the difference between cooled and uncooled detectors.

## Slide 19
MOD 2 • [INSTRUMENTATION] • SLIDE 18/30
Stray Light & Optical Cross-Talk
CORE TAKEAWAY: Scattered photons reflecting off internal spectrometer walls contaminate adjacent detector pixels, degrading spectral purity.
Stray light: Radiation of unwanted wavelength reaching a detector channel due to imperfect grating diffraction or wall scatter.
Most severe in deep absorption bands (e.g. O₂ A-band at 760 nm or water bands) where true signal is near zero.
Corrected using specialized stray light matrices calibrated with laser emission lines or cutoff filters.
References: Zong, Y. et al. (2006). Simple method for stray-light correction on array spectroradiometers. Applied Optics, 45(6), 1111-1119.
Notes: Explain how stray light leads to underestimation of deep absorption feature depths if not mathematically calibrated.

## Slide 20
MOD 2 • [QUALITYASSURANCE] • SLIDE 19/30
Temperature Sensitivity and Thermal Drift
CORE TAKEAWAY: Field ambient temperatures fluctuating from 15°C morning to 35°C afternoon induce internal mechanical and optical shifts.
Thermal expansion of optical bench alters diffraction grating angles and focal plane alignment (sub-nanometer wavelength drift).
InGaAs detector responsivity shifts with temperature, especially near bandgap cutoff boundaries.
Mitigation: Continuous monitoring of internal instrument thermistors and frequent white reference / dark current re-acquisitions.
FIELD PROTOCOL NOTE
Keep the spectrometer body in the shade when operating in hot desert or summer environments.
References: MacArthur, A. (2007). A field procedure to compensate for temperature drift in portable spectrometers. NERC FSF.
Notes: Highlight the built-in temperature logging in modern SVC and ASD instruments.

## Slide 21
MOD 2 • [INSTRUMENTATION] • SLIDE 20/30
Single-Beam vs. Dual-Beam Field Architecture
CORE TAKEAWAY: Dual-beam instruments record target radiance and downwelling irradiance simultaneously, canceling temporal solar fluctuations.
Single-beam (Standard ASD/SVC): One optical head alternates sequentially between White Reference panel and target.
Dual-beam (e.g. Dual ASD, PP-Systems, Unispec-DC): Two intercalibrated channels measure white reference / cosine diffuser and target at the exact same millisecond.
Dual-beam eliminates error from moving clouds, but introduces complex inter-channel sensor cross-calibration challenges.
FIELD PROTOCOL NOTE
In dual-beam setups, exchange fiber heads periodically over the reference panel to track inter-channel drift.
References: Milton, E. J. et al. (2009). Progress in field spectroscopy. Remote Sensing of Environment.
Notes: Compare the operational trade-offs: single-beam is simpler and lighter; dual-beam is essential in variable temperate climates like the UK.

## Slide 22
MOD 2 • [PROTOCOL] • SLIDE 21/30
Pistol Grips, Mounting Rigs and Tripods
CORE TAKEAWAY: Rigid mechanical mounting eliminates operator hand-tremor, maintains nadir viewing geometry, and sets fixed sensor heights.
Handheld pistol grip: Convenient for rapid surveys, but prone to operator zenith tilt and operator shadow artifacts.
Counterbalanced boom tripod: Holds foreoptic 2 to 3 meters away from the operator and metal legs, minimizing adjacency shadows.
Integrated bubble levels and laser pointers: Ensure precise nadir alignment (θ_r = 0°) and target center verification.
FIELD PROTOCOL NOTE
Always check the bubble level before triggering every scan. A 10° tilt completely invalidates BRDF assumptions.
References: NERC FSF (2020). Operations Manual.
Notes: Show photos of well-designed field boom arms versus sloppy handheld operator practices.

## Slide 23
MOD 2 • [QUALITYASSURANCE] • SLIDE 22/30
Spectralon® White Reference Reflectance Standards
CORE TAKEAWAY: Sintered polytetrafluoroethylene (PTFE) provides > 99% diffuse Lambertian reflectance from 350 to 1600 nm (> 95% to 2500 nm).
Manufactured by Labsphere: Porous, sintered thermoplastic with isotropic bulk volume scattering.
High thermal and chemical stability; impervious to water (hydrophobic).
Certified by NIST with a calibration certificate listing exact spectral reflectance factors at 50 nm intervals.
FIELD PROTOCOL NOTE
Never touch the active white surface of a Spectralon panel with bare fingers or let dust accumulate.
References: Labsphere Inc. (2018). Spectralon Diffuse Reflectance Standards and Targets Technical Guide.
Notes: Explain the microscopic porous structure of PTFE that creates almost perfect diffuse scattering through multiple internal refractions.

## Slide 24
MOD 2 • [QUALITYASSURANCE] • SLIDE 23/30
White Reference Degradation and Contamination
CORE TAKEAWAY: Panel contamination by skin oils, dust, and UV photochemical yellowing degrades short-wavelength reflectance.
Hydrocarbon contamination (fingerprints, volatile organics) creates deep C-H absorption bands in the SWIR (1200, 1720, 2300 nm).
UV solarization causes yellowing, dropping blue/UV reflectance from 99% down to 92%.
A contaminated reference panel makes targets appear artificially high in reflectance (apparent reflectance > 100%).
FIELD PROTOCOL NOTE
Store reference panels in clean desiccated cases and clean annually with waterproof emery sandpaper under deionized water.
References: MacArthur, A. et al. (2012). The calibration and characterization of field spectrometers. NERC FSF.
Notes: Demonstrate how to re-sand and clean a Spectralon panel according to Labsphere instructions.

## Slide 25
MOD 2 • [INSTRUMENTATION] • SLIDE 24/30
Spectral Transmittance & Absorbance Measurement Modes
CORE TAKEAWAY: Spectroradiometers can measure transmittance (T) and absorbance (A) in addition to reflectance using specialized accessories.
Transmittance: Attaching an integrating sphere allows measurement of total directional-hemispherical transmittance of single leaves.
Absorbance: A = 1 - R - T (by conservation of energy in non-scattering systems).
Essential for calibrating leaf radiative transfer models like PROSPECT which require both leaf reflectance (ρ) and transmittance (τ).
MATHEMATICAL FORMULATION
A(\lambda) + R(\lambda) + T(\lambda) = 1.0
First law of thermodynamics applied to radiant energy conservation.
References: Jacquemoud, S., & Ustin, S. L. (2019). Leaf Optical Properties. Cambridge University Press.
Notes: Explain why measuring only leaf reflectance is insufficient to parameterize PROSPECT; leaf transmittance is equally critical.

## Slide 26
MOD 2 • [INSTRUMENTATION] • SLIDE 25/30
Integrating Spheres in Field Spectroscopy
CORE TAKEAWAY: Integrating spheres collect all scattered light across the entire 4π or 2π hemisphere, enabling pure hemispherical measurements.
Hollow sphere coated internally with highly reflective diffuse Spectralon or Zenith coating.
Multiple internal diffuse reflections homogenize radiation uniformly over the inner wall area.
Measures directional-hemispherical reflectance factor (DHRF) and directional-hemispherical transmittance factor (DHTF).
References: Labsphere Inc. (2015). Integrating Sphere Theory and Applications.
Notes: Discuss field-portable integrating spheres (e.g. ASD RTS-3ZC, SVC LC-RP Pro).

## Slide 27
MOD 2 • [PROCESSING] • SLIDE 26/30
Calibration Files & Raw Digital Numbers to Radiance
CORE TAKEAWAY: Conversion from raw digital numbers to physical spectral radiance requires the instrument calibration coefficient file.
Each instrument possesses a factory-calibrated responsivity vector: R_cal(λ) [DN / (W m⁻² sr⁻¹ nm⁻¹)].
Radiance: L(λ) = (DN(λ) - DN_dark(λ)) / (t_int · R_cal(λ)).
When computing relative reflectance against a white reference, R_cal(λ) cancels out in the ratio, which is why reflectance is robust against absolute calibration drift.
MATHEMATICAL FORMULATION
\rho(\lambda) = \frac{L_{\text{target}}(\lambda)}{L_{\text{ref}}(\lambda)} \cdot \rho_{\text{ref}}(\lambda) = \frac{\frac{\text{DN}_{\text{t}} - \text{DN}_{\text{d}}}{t_{\text{int,t}} \cdot R_{\text{cal}}}}{\frac{\text{DN}_{\text{r}} - \text{DN}_{\text{d}}}{t_{\text{int,r}} \cdot R_{\text{cal}}}} \cdot \rho_{\text{ref}}(\lambda)
Cancellation of absolute responsivity R_cal in the reflectance ratio calculation.
References: Milton, E. J. (1987). International Journal of Remote Sensing.
Notes: Point out why reflectance can be measured accurately in the field even if absolute radiometric calibration has decayed by 10%.

## Slide 28
MOD 2 • [QUALITYASSURANCE] • SLIDE 27/30
Power Systems, Battery Care & Thermal Management
CORE TAKEAWAY: Stable battery voltage is crucial to avoid baseline electronic shifts in Peltier cooling and InGaAs detector amplifiers.
Lithium iron phosphate (LiFePO4) or high-capacity lithium-ion field battery packs.
Voltage drops below operating threshold induce subtle noise spikes and premature instrument shutdown.
Always carry at least two fully charged spare batteries and an in-field 12V DC vehicle charging adapter.
FIELD PROTOCOL NOTE
Cold weather (< 5°C) reduces lithium battery capacity by up to 40%; keep spare batteries inside thermal jackets.
References: NERC FSF (2020). Operational Field Guide.
Notes: Share field horror stories of spending two hours hiking to a remote field site only to find dead batteries.

## Slide 29
MOD 2 • [INSTRUMENTATION] • SLIDE 28/30
Software Interfaces: RS3, SVC PC, and SAMS
CORE TAKEAWAY: Instrument control software handles auto-integration, dark current acquisition, spectral logging, and real-time QA displays.
ASD RS3 / Indico Pro: Real-time graphical display of DN, radiance, and reflectance; automated splice correction.
SVC PC Software: Manages Bluetooth telemetry, automated camera integration, and internal file conversions.
SAMS (Simple Automated Measurement Software): Open-source automated field spectroscopy collection suite.
References: Analytical Spectral Devices (2012). RS3 User Manual.
Notes: Show the live software GUI and highlight the key indicators students must watch during active data capture.

## Slide 30
MOD 2 • [CASESTUDY] • SLIDE 29/30
Case Study: Point Spectrometry in Mineral Exploration
CORE TAKEAWAY: Point spectroscopy in the Atacama desert demonstrates rapid mineral mapping of hydrothermal alteration zones.
Hydrothermal porphyry copper deposits exhibit concentric alteration halos: potassic, phyllic, argillic, propylitic.
Field spectroradiometers identify white mica Al-OH composition shifts (2198 nm sodic paragonite vs 2206 nm potassic muscovite vs 2220 nm phengite).
Wavelength shifts correlate directly with paleofluid temperatures and proximity to ore bodies.
References: Thompson, A. J. B. et al. (1999). Alteration mapping in exploration: application of short-wave infrared (SWIR) spectroscopy. SEG Newsletter.
Notes: Present the geological cross-section and show how 500 point spectra mapped the entire hydrothermal alteration gradient.

## Slide 31
MOD 2 • [INSTRUMENTATION] • SLIDE 30/30
Module 2 Summary & Transition to Calibration Protocols
CORE TAKEAWAY: Hardware architecture dictates data quality: numerical aperture, gratings, detector cooling, and ADC dynamic range.
Module 2 reviewed: Optical layout, fiber cables, silicon vs InGaAs detectors, Spectralon white references, and SNR optimization.
Next in Module 3: Laboratory calibration, wavelength standards (Mercury-Argon), radiometric calibration with integrating spheres, and uncertainty budgets.
Laboratory practical: Processing raw binary spectroradiometer files, removing detector steps, and computing SNR curves.
References: Milton, E. J. et al. (2009). Remote Sensing of Environment, 113, S92-S109.
Notes: Wrap up hardware lecture, confirm understanding of detector bandgap limits and fiber care.